## Slide 1: Lab-based phase-contrast micro-CT for osteochondral tissue at ~1 µm resolution
Álvaro José González Grajales
11/06/2026
Notes: Cover title slide using a full UCL Dark Purple colour as the background. The slide layout includes a main heading in the top left corner. The presenters’ names and the date are shown in the bottom left corner. The UCL logo is positioned in the bottom right corner of the slide.

## Slide 2: Introduction
Osteochondral tissue (OC): Load-bearing composite structure found in joints comprising cartilage, Calcified cartilage and the underlying subchondral bone
Often subjected to trauma or degradation leading to osteoarthritis ([~303.1 million prevalent cases])
The limited regenerative capacity of avascular, aneural and alymphatic cartilage poses a major challenge after injury or disease.
In its natural state, OC tissue presents well-defined cellular and extracellular matrix architecture that has not been successfully restored using conventional treatment
(Niu et al., 2023)
(Pueyo Moliner et al., 2025)
(Davis S.L., 2025)
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 3: Need & aim
The study of OC tissue has largely relied on 2D imaging techniques that require destructive sample preparation (Histology is “gold standard”)
2D imaging fails to capture the relation between cell shape, size and distribution with the maintenance of the ECM
Volumetric, non-destructive techniques have emerged, mainly MRI and X-ray CT, for the assessment of natural or engineered tissue.
However, a lack of contrast (cartilage) and a limited resolution (mm-cm) have limited its application for the OC tissue.
Use lab-based phase contrast µCT to develop a non-destructive imaging methodology that provides a reliable and easily accessible assessment tool for natural or engineered OC tissue with sufficient resolution and contrast to visualize cells within the extracellular matrix in 3D
(Nieminen H.J, et. al., 2017)
(Guermazi, A., Eckstein, F., Gold, G. et al.., 2026)
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 4: Inline phase-contrast imaging
In contrast to attenuation-based imaging, phase-contrast exploits the phase-shift experienced by the X-rays as they interact with the sample to generate imaging contrast
Leveraging the complex refractive index to generate contrast in the image offers an opportunity to increase SNR, especially when dealing with weakly attenuating samples
For propagation-based PC: the refraction process induces an intensity modulation pattern that corresponds to a mixture between δ and β
To extract phase and amplitude information from the image a phase-retrieval algorithm must be applied: Paganin’s single-image method (Non-quantitative)
The phase-retrieved projections can be reconstructed into 3D volumes to describe internal structure of the object with proper tomographic reconstruction method
(Paganin, et.al, 2002)
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 5: µCT at UCL (NXCT)
Detector pixel size: 4.5μm (3200x2200)
Optical magnification | Effective pixel size | FOV
5x | 0.9μm | ~2mm
10x | 0.45μm | ~1mm
20x | 0.225μm | ~0.5mm
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 6: FIRST STAGE: is it possible?
Could the LB-XPCi-uCT effectively be used for assessing OC tissue?

## Slide 7: Modelling the sample
(Pueyo Moliner et al., 2025)
As rule-of-thumb one should aim for 25-75% of X-ray transmission
Cartilage
Bone
Linear attenuation coefficient
“Rule of mixtures” kind of law
%O | %C | %H | %N | %P | %S | %Na | %Cl
74 | 9.9 | 9.6 | 2.2 | 2.2 | 0.9 | 0.5 | 0.3
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 8: Modelling the sample
Only for a 1 mm thick sample both the cartilage and bone showed transmission ~[25%-75%] for either Cu or Mo
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 9: Resolution test
Could the system resolve ~10μm features (cell)
Calculate the MTF
2.5um
1 um
0.5 um
1.5um
0.75 um
Image quality does not only depend on the detector…but also on the stability of the system
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 10: Resolution test
Landed
Floating
Elastomer
For both the floating and elastomer configurations the system exhibits an MTF50 at 500 LP/mm, corresponding to an effective spatial resolution of <1 µm
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 11: Second STAGE: Pilot study
Can LB-XPCi-uCT effectively visualize individual chondrocytes (~10um) within cartilage

## Slide 12: Methods
Sample preparation protocol
Acquisition
Image optimization: Phase retrieval & CT reconstruction
#projections?
Exposure time?
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 13: Results: Image optimization and reconstruction
COR and ring artefact correction were optimized to improve the visualization of cells without inducing distortion of the image
Optimisation of the δ/β ratio at the cell-ECM boundary enhances cell visibility and improves noise suppression
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 14: Results: Image optimization and reconstruction
SNR and CNR improve with the integration time following a relation (Poisson statistics)
Even at 10s exposure the cellular structures were observed reducing imaging time to ~4hours
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 15: Results: General Assessment
Mean chondrocyte diameter measured 10.34 ±1.84 μm is consistent with literature (~10μm)
Structural features are comparable between imaging modalities
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 16: Results: Volumetric segmentation
Pixel-wise 3D volume segmentation
Feature selection in the classification model (Intensity, texture, edges)
Ilastik
Visual examination shows good agreement between chondrocyte segmentation and images.
Morphological examination and quantitative information has yet to be obtained
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 17: Conclusions
Can LB-XPCi-uCT effectively visualize individual chondrocytes (~10um) within osteochondral tissue/scaffolds?
Yes, it can! It is even possible to segment the cells throughout the volume opening the possibility for quantitative assessments
The volume showed comparable morphological features to histological imaging
BUT…
Although not “destructive” the sample preparation protocol followed still required fixation/dehydration which may change both the morphological and mechanical characteristics of the OC tissue…
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 18: Third stage: let’s challenge the methodology
Is it possible to visualize the same features in non-fixed & hydrated tissue resembling physiological conditions?

## Slide 19: Imaging hydrated samples
~ 16000 projections * 10sec =~ 44 hours (non-ideal)
Although this protocol has not yet been optimized, we are still able to observe cellular features in non-fixed, non-stained, hydrated samples
Opens the possibility of evaluating what are the implications of sample preparation protocols and assess the morphological characteristics of cells in a near-physiological state
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 20: Proposal: Sample preparation and Correlative validation
Implement a “Daisy chain” registration strategy to evaluate morphological changes in the sample during destructive sample preparation protocols
Correlative analysis between XPC-CT-histology-EM
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 21: Conclusion
Lab-based Phase-contrast µCT can visualize individual chondrocytes (cells) in 3D
Cell segmentation is possible without destructive histology
Imaging hydrated tissue in the lab is very challenging but feasible
Correlative imaging will help quantify sample preparation induced changes
Future work
Quantitative Morphological Analysis: Finalize the extraction of quantitative metrics (shape, size, and distribution) from segmented 3D volumes to move beyond qualitative visual examination
"Daisy Chain" Correlative Validation: Complete the registration strategy to correlate XPC-µCT data with histology and electron microscopy, allowing for a precise evaluation of morphological changes induced by traditional sample preparation
Notes: This content slide features a headline at the top left, followed by two columns, each providing space for an image with accompanying copy beneath.

## Slide 22
Gracias